Autoclave

SM300/500/510

Inspection Procedure

Yamato Scientific Co.,Ltd.

Regilar inspection

Small size pressure vessels inspection

Inspection Procedure

| Preparation | －Necessary equipments for this inspection－ 1)Tools Phillips Head Screwdriver, Wrench(size: 10, 14, 17, 19, 24, and 30), Adjustable Wrench, Hex Wrench(size: 2.5) 2)The calibrated Temperature Sensor(measured up to 200℃) and the Temperature Recorder 3)The calibrated Pressure Gauge: Range up to 0.3MPa 4)Pressure Device(Compressor<up to 0.22MPa>, or High pressure Inert gas container, and or Bicycle Pump) 5)Pressure-proof Rubber Hose (ID 10mm), Hose band, and fitting |
| --- | --- |
| 1 Appearance Inspection | Check the chamber and the lid has no scratches, no stain and no deformations. ●Remove the scratches and the stain with abrasives. Check the arm surface has no scratches/no crakes. Require to replace the arm, if the scratches/the crakes will be found. Check the handle thread has no scratches, no stain and the handle rotates smoothly. Apply the grease such as Mitsubishi multi-purpose grease No.0 or the equivalent, if the thread has stain and/or its rotation is not smoothly. Check the arm nut does not turn at the lid open. The arm nut will be loosened when it is turn. Two set screws must be removed and then the arm nut be tightened by hand. Set two set screws very tight. Check the gasket has no scratches and no deformation. Replace the gasket, if scratches and/or deformation is found. |

Regilar inspection

Small size pressure vessels inspection

Inspection Procedure

| 2 Remove the covers | Remove the covers in order of the rear one, the right one, the left one. |
| --- | --- |
| 3 Piping line inspection | Check the drain valve, the drain solenoid valve and the exhaust solenoid valve will not be loosened and/or wetted at their connections. ●Please tighten those valves, if the above phenomena will be observed. Check all fitting connections will not be loosened and/or wetted at them. Please tighten those valves, if the above phenomena will be observed. Check all pipe lines will not be wetted. ●Require to replace the pipe line, if the above phenomena will be observed. |

Regilar inspection

Small size pressure vessels inspection

Inspection Procedure

| 4 Pressure Gauge Inspection | Please connect each devices according as above configuration. When the chamber is pressurized set at 0.05MPa, 0.1MPa and 0.15MPa, the indicating gauge pressure on the autoclave built-in gauge should be compared with the setting pressure. It will be passed, if the pressure reading on the autoclave built-in gauge is within ±0.005MPa of the set pressure. However, need to replace the autoclave built-in gauge, if the pressure reading will be out of its tolerance. |
| --- | --- |
| 5 Safety Valve Inspecction | To check the working pressure of the safety valve, increase the pressure starting at 0.15MPa, gradually. It is normal, if the working pressure read between 0.18MPa thru. ～0.2MPa. Do not pressurize more than 0.22MPa, even if the valve will not activated. If the working pressure will not in the range of 0.18MPa thru. 0.2MPa, then the working pressure need to adjust the following procedure. Loosen the stable thread. Adjust the working pressure by the adjusting thread. Tighten the stable thread after the pressure adjustment. |

Regilar inspection

Small size pressure vessels inspection

Inspection Procedure

| 6 Solenoid Valve, Drain Valve Leakage Inspection | Please close the drain valve right away when the safety valve is activated. |
| --- | --- |
|  | Block the holes on the exhaust hose by the electrical tap. The silence is submerged under the water of the container such as a cap see if the air bubble is come out or not. In case of coming out the air bubble: One of the exhaust solenoid valve, or the drain solenoid valve, or the drain valve have the defect for the air leakage, so that the repair or the valve replacement will be required. In case of coming out no air bubble but decreasing the chamber pressure: There may be the air leakage on the pressured chamber or the pipe line. Please find out the leakage portion by the soapsuds and fix the air leakage properly by the re-tightening and the replace parts and so forth. |

Regilar inspection

Small size pressure vessels inspection

Inspection Procedure

| 7 Temperature Accuracy | The temperature accuracy will be checked by one of the following method as below. A calibrated temperature sensor(with a temperature recorder) or a fixed-point thermometer will be installed around the chamber control sensor. The autoclave will be operated by the sterilizing cycle set at 128℃. The displayed temperature(128℃) on the autoclave control panel is checked against the temperature that should be indicated between 128℃ thru. 130℃. Note: A temperature sensor or a fixed-point thermometer should be installed to the service port(R1/4-female) The autoclave will be operated and the displayed gauge pressure and the displayed temperature will be checked as the below table. Please refer the relation between the temperature and the pressure in the chamber at the ambient pressure that is in 1,013 hPa. 120℃…0.097MPa 121℃…0.104MPa 122℃…0.110MPa 123℃…0.117MPa 124℃…0.124MPa 125℃…0.131MPa 126℃…0.138MPa 127℃…0.145MPa 128℃…0.153MPa 129℃…0.161MPa 130℃…0.169MPa |
| --- | --- |

Through the regular inspection above, you have any difficulties on the repair or the replacement of the parts, please contact the local Yamato agent or Yamato Scientific sales Office.

- Regilar inspection

Small size pressure vessels inspection

－Inspection Items and Record Table－

| Inspection Date Inspection Items | Month/Date/Year | Month/Date/Year | Month/Date/Year | Month/Date/Year |
| --- | --- | --- | --- | --- |
| Chamber | Scratch, Stain Deformation |  |  |  |  |
| Lid | Scratch, Stain Deformation |  |  |  |  |
| Arm | Scratch（Crakes） |  |  |  |  |
| Handle Thread | Scratch, Stain |  |  |  |  |
| Arm Nut | Loosened Nut |  |  |  |  |
| Gasket | Scratch Deformation |  |  |  |  |
| Drain Valve | Water Leakage |  |  |  |  |
| Drain Solenoid valve | Water Leakage |  |  |  |  |
| Exhaust Solenoid Valve | Water Leakage |  |  |  |  |
| Safety Valve | Working Pressure |  |  |  |  |
| Pipe line & Fittings | Scratch Loosened fitting Water leakage |  |  |  |  |
| Pressure Gauge | Pressure Accuracy |  |  |  |  |
| Temperature Accuracy | temperature Accuracy |  |  |  |  |
| Checked by |  |  |  |  |
| Proved by |  |  |  |  |

・If no problem are found, then mark “(” in the box. If any problems are found, describe their phenomena and how to fix them.

# Text boxes

④アーム用ナット

③Screw

Handle

②Arm

- Lid

⑤Gasket

(Behind the Lid)

Set

Screws

- Chamber

④Arm Nut

Rear Cover

Left Cover

Right Cover

Install the hose band

Close the lid

Calibrated Pressure Gauge

Pressure

Device

Stable

Thread

Adjusting Thread

Coil Spring

【Disassembly Drawing】

Working Pressure

Down

Working Pressure Up

Loosen Dir.

Tighten

Dir.

Pressure

Device

Close

this valve

Block

these holes